# Diffusion MRI toolkits still disagree on how to weight a tensor fit, a decade after the question was settled

Busra Mutlu¹*

¹ Department of Neuroimaging, King’s College London, London, United Kingdom

* Corresponding author. Busra Mutlu, Department of Neuroimaging, King’s College London, De Crespigny Park, London SE5 8AF, United Kingdom. E-mail: [KCL EMAIL]

## Abstract

Background. Veraart et al. (2013) showed that weighting a linear tensor fit by the measured signal degrades accuracy, and that weights from a predicted signal perform better. Whether the major toolkits apply that result has not been examined.

New method. We compared tensor fits from FSL, MRtrix3 and DIPY on two open datasets with every input held constant, varied the estimator rather than the toolkit, and measured accuracy against phantoms at SNR 30, 20 and 10.

Results. The toolkits do not agree, and none of it is visible in the command a user runs: MRtrix3 applies the recommended multi-step weighting by default and DIPY a two-pass equivalent, while FSL performs ordinary least squares unless asked otherwise and its --wls uses the scheme advised against. Implementation is not the variable — FSL --wls and MRtrix3 on the same estimator agree exactly (r = 1.0000, MAE 0.0000). At SNR 30 the two weighting families differ negligibly; at SNR 10 the measured-signal family underestimates FA by 0.0177 and the predicted-signal family overestimates it by 0.0111. On real data the difference reached 0.26 standard deviations of white matter FA, more than replacing one toolkit with another.

Comparison with existing methods. Earlier work reports pipeline variability as software-dependent. Attributing it to the estimator and to differing defaults locates it precisely and makes it correctable.

Conclusions. The estimation question has an answer in the literature; the toolkits have not converged on it. Users cannot see which scheme they apply, and at routine SNR the choice is not negligible. Methods sections should report the estimator and its settings, not only the toolkit.

Keywords: diffusion MRI; tensor fitting; weighted least squares; estimator accuracy; reproducibility; software defaults

## Introduction

Diffusion magnetic resonance imaging (dMRI) is the principal non-invasive technique for probing white matter microstructure and structural connectivity in the living human brain (Basser et al., 1994; Jones, 2010). By encoding the directional displacement of water molecules, dMRI provides access to biophysical indices — including fibre orientation, axon density, and myelin integrity — that are invisible to conventional anatomical MRI (Beaulieu, 2002). Applications span fundamental neuroscientific questions about brain organisation and inter-individual variability, as well as clinical contexts including the characterisation of white matter alterations in neurodegenerative diseases, psychiatric conditions, and neurodevelopmental disorders (Catani and Thiebaut de Schotten, 2008; Jones, 2010).

The dMRI analysis ecosystem is dominated by three major software packages. FSL (FMRIB Software Library), developed at the University of Oxford, provides robust preprocessing, DTI fitting, and voxelwise group analysis via the Tract-Based Spatial Statistics (TBSS) pipeline (Smith et al., 2004; Jenkinson et al., 2012). MRtrix3, developed at The Florey Institute of Neuroscience and Mental Health, specialises in constrained spherical deconvolution (CSD) and high-fidelity fibre orientation modelling (Tournier et al., 2019). DIPY (Diffusion Imaging in Python) offers a flexible, Python-native implementation of the full dMRI pipeline and is widely used when algorithmic transparency or custom extensions are required (Garyfallidis et al., 2014).

Despite their complementary strengths, these toolkits present a significant practical barrier. Each employs distinct command-line syntax, file format conventions, and — in several cases — different terminology for the same operation. For example, brain extraction is called bet in FSL, dwi2mask in MRtrix3, and median_otsu in DIPY. Mean diffusivity is reported as MD in FSL, ADC in MRtrix3, and md in DIPY. These surface-level differences cause genuine confusion for researchers trained in one ecosystem who attempt to read, reproduce, or extend work conducted in another.

A more substantive concern is methodological reproducibility. Published dMRI studies rarely justify their choice of software, and it is not always clear whether reported group differences in white matter metrics reflect genuine biological effects or implementation-level differences between tools. Software-specific choices — brain masking algorithm, denoising method, tensor fitting algorithm — can introduce non-trivial variability in derived metrics such as fractional anisotropy (FA) and mean diffusivity (MD), and this has been demonstrated at the level of whole pipelines (Bhagwat et al., 2021; Richie-Halford et al., 2022).

Such studies leave a narrower question open: when the only thing that differs is the software performing the tensor fit, how much do the toolkits disagree, and what inside the software is responsible? Answering the first part requires holding every other input fixed. Answering the second requires going further, and manipulating the candidate cause directly.

We built a containerised environment in which all three toolkits run on byte-identical data, used it to run the controlled comparison, and then followed the difference to its source. The instrument and the measurement are reported together because neither is much use alone.

Four questions are addressed. First, how far do FSL, MRtrix3 and DIPY agree on fractional anisotropy and mean diffusivity given identical input? Second, are the toolkits in fact running the same estimator, or does each apply a different one by default? Third, if the estimator is matched, does the disagreement persist? Fourth, when the toolkits disagree, is any of them wrong — a question that requires a ground truth the real data cannot supply.

The estimation question itself is not open. Veraart et al. (2013) compared weighting a linear tensor fit by the squares of the measured diffusion-weighted signals against weighting by the squares of a predicted signal reconstructed from an earlier parameter estimate. They found the negative effect of the former on accuracy to be, in their words, surprisingly high, and reported that multi-step weighting performed better and in some cases outperformed non-linear least squares. Koay et al. (2006) had earlier set out the general framework for least-squares estimation in DTI.

What has not been examined is whether the toolkits the field actually uses apply that result. Each of the choices we identify is documented by the project that made it: MRtrix3 states that dwi2tensor iterates its reweighting twice, DIPY cites Chung et al. (2006) for its weights, and FSL’s --wls flag appears in its own help text. Reading three manuals establishes that the components differ. It does not establish which toolkit implements the recommendation, what the difference costs at the signal-to-noise ratios of routine acquisition, or how it compares with the effect of changing toolkit — and those are the questions facing someone choosing how to fit a tensor.

We report the comparison on two open datasets acquired at different sites under different protocols, on a synthetic phantom whose generating eigenvalues are known, and under a direct manipulation of the weighting scheme. Every value can be regenerated from openly available data.

## Materials and Methods

### Data

All demonstrations use the Stanford HARDI dataset distributed by the DIPY project under an open-access licence (Rokem et al., 2015). The dataset comprises 160 volumes (10 b = 0, 150 DWI at b = 2000 s/mm²) with voxel size 2 × 2 × 2 mm and matrix 81 × 106 × 76. It is downloaded automatically on first launch via DIPY’s get_fnames("stanford_hardi") utility without credentials. A brain mask is generated at download time using median_otsu and stored alongside the raw NIfTI data.

The quantitative benchmark reported below additionally uses the Sherbrooke 3-shell dataset, distributed openly as part of the DIPY sample data collection (Garyfallidis et al., 2014) and retrieved via get_fnames("sherbrooke_3shell"); like the Stanford data it requires no credentials. It comprises 193 volumes (1 b = 0; 64 directions at each of b = 1000, 2000, and 3500 s/mm²) with voxel size 2 × 2 × 2 mm and matrix 128 × 128 × 60. The dataset is distributed with DIPY rather than accompanied by a dedicated acquisition publication, and we use it here as an independent test acquisition rather than as a characterised reference sample. Because it differs from the Stanford data in site, protocol, angular sampling, b-value range, and — importantly for the results below — the number of b = 0 volumes, analysing both provides an independent replication of the inter-tool comparison rather than a second look at one acquisition.

The application also supports Human Connectome Project (HCP; Van Essen et al., 2013) multi-shell data (b = 1000/2000/3000 s/mm², 90 directions per shell) for users with HCP data access credentials, enabling demonstration of multi-tissue CSD which requires multiple non-zero b-value shells.

### Comparison design

Beyond the interactive pipeline, we conducted a systematic quantitative comparison of DTI scalar metrics and brain masks across all three toolkits, on both the Stanford HARDI and the Sherbrooke 3-shell datasets. This analysis serves two purposes: it validates that the platform produces outputs consistent with established diffusion MRI principles, and it constitutes a reproducible, openly available benchmark of inter-tool agreement under controlled conditions — identical input volumes, identical gradient tables, and a single shared brain mask.

The design is summarised in Figure 1. It comprises two comparisons that make opposite choices about the input, and the distinction between them matters for interpreting the results. For brain extraction, each tool is run on the input its algorithm expects, so the comparison reflects the difference a user would actually encounter. For tensor fitting, every tool is given byte-identical input, so the comparison isolates the fitting stage from every other source of variation.

Brain mask agreement was assessed using the Dice similarity coefficient (DSC) between each pair of binary masks produced by FSL bet, MRtrix3 dwi2mask, and DIPY median_otsu. Each tool was run in its own idiomatic configuration rather than forced onto a common input, because the input a tool is designed to consume is part of the algorithm being compared: bet operates on the mean b = 0 image, dwi2mask on the full DWI series, and median_otsu on the full series with the b = 0 volumes indexed. The DSC therefore quantifies the end-to-end difference between the three brain-extraction routines as a user would actually invoke them. DSC values are reported in Table 2.

Separately, and to keep brain extraction from confounding the metric comparison, a single shared brain mask generated by median_otsu was supplied to all three tensor fits. A white matter reference mask, defined as voxels where all three tools jointly yielded FA > 0.2, was used for all subsequent metric comparisons.

DTI metric agreement was assessed for FA and MD on both datasets. To isolate tensor-fitting differences from every other source of variability, all three tools were applied to the identical input: the same unprocessed volumes, the same bvals and bvecs, and the shared median_otsu brain mask described above. No denoising or eddy-current correction was applied before fitting. This omission is deliberate rather than a shortcut: any preprocessing step would itself be a toolkit-specific choice, and applying one tool’s preprocessing to all three would confound the very comparison it is meant to enable. The consequence, which we state plainly, is that the reported agreement characterises the tensor-fitting stage alone and not a complete analysis pipeline.

For the multi-shell Sherbrooke data a further step was required. The single-tensor model assumes monoexponential signal decay, which does not hold across shells; FSL dtifit and MRtrix3 dwi2tensor silently fit every volume they are given, whereas DIPY requires the shell to be selected by the user. Left unaddressed, the three toolkits would have fitted different data and any difference between them would confound tensor fitting with shell selection. We therefore extracted the b = 0 and b = 1000 s/mm² volumes (65 of 193) once, and passed that identical subset to all three toolkits. The Stanford data is single-shell, so all 160 volumes were used.

Voxelwise Pearson correlation coefficients (r), Spearman rank correlations (ρ), and mean absolute error (MAE) were computed over white matter voxels, and Bland–Altman analysis was performed for each cross-tool pair to characterise the distribution of voxelwise differences and identify any systematic bias.

Statistics were restricted to voxels whose values are physically admissible in both tools of a given pair: FA within [0, 1], and MD positive and no greater than the diffusivity of free water at body temperature (3.0 × 10⁻³ mm²/s). This restriction is necessary rather than cosmetic. Unconstrained linear tensor fitting can return negative eigenvalues, which drive FA above unity and MD below zero; because Pearson correlation is dominated by extreme values, a few hundred such voxels are sufficient to depress the apparent MD agreement from r = 0.997 to r = 0.118 while changing the mean absolute error by less than 2%. The per-tool counts of these non-physical voxels are reported in the Results as a substantive inter-tool difference in their own right.

Results are reported in Tables 3 and 4 and in Figures 3 and 4, with the corresponding Sherbrooke figures provided as Supplementary Figures S1 and S2. All analyses were performed using the companion scripts scripts/generate_fa_maps.py and scripts/compute_fa_comparison.py distributed with the repository. Every value reported here is reproducible from openly available data with no credentials, by running:

python scripts/generate_fa_maps.py      --subject stanford
python scripts/compute_fa_comparison.py --subject stanford
python scripts/generate_fa_maps.py      --subject sherbrooke --shell 1000
python scripts/compute_fa_comparison.py --subject sherbrooke

### Matching the estimator

The comparison above compares toolkits at their usual settings, which is what a user encounters, but it does not guarantee that the same estimator is being applied. MRtrix3 dwi2tensor performs weighted least squares followed by two iterations of reweighting by default. To place all three on plain WLS we repeated the comparison with dwi2tensor -iter 0, which stops after the first weighted fit, alongside FSL dtifit --wls and DIPY fit_method="WLS". Both configurations are reported: the default arms because they describe practice, the matched arms because they isolate the estimator.

### Manipulating the weighting scheme

MRtrix3 documents its first stage as weighting by the empirical signal intensities; DIPY documents its weights as the squared signal predicted by an initial ordinary least-squares fit, following Chung et al. (2006). To test whether this accounts for the residual difference we used DIPY’s facility for user-supplied weights, running the same DIPY code twice on the same data — once with its default predicted-signal weights, once with the measured signal — and compared each against the FSL fit. Signals were floored at unity before both fits so that the weighted design matrix stays non-singular where a voxel reads zero; both arms received identical clipped data.

### Accuracy against a known ground truth

Agreement between toolkits cannot say whether any of them is accurate. We therefore fitted synthetic phantoms whose generating eigenvalues are known, at three noise levels. A single noise level would not settle the question: the theoretical objection to measured-signal weighting is that the weights are correlated with the noise they are meant to downweight, so any difference should be small at high SNR and grow as SNR falls. Phantoms were generated at SNR 30, 20 and 10 at b = 0, spanning from a favourable acquisition to one typical of high b-value shells. Each phantom contains: an isotropic region with eigenvalues 0.9, 0.9, 0.9 × 10⁻³ mm²/s (FA = 0, MD = 0.90 µm²/ms) and a single-fibre region with 1.4, 0.35, 0.35 × 10⁻³ mm²/s (FA = 0.7071, MD = 0.70 µm²/ms), A crossing-fibre region is present in the phantom but excluded from scoring, since no single tensor is correct there by construction. Voxels within two of a region boundary were also excluded, as they mix tissue types. Bias and root-mean-square error against truth are reported for each arm.

### Sensitivity to preprocessing

The comparison applies no denoising or eddy-current correction, so that the fit is examined at a defined starting point rather than after an arbitrary preprocessing choice. To test whether the findings survive realistic preprocessing, we repeated the whole comparison on MP-PCA denoised data. Denoising was performed once, with MRtrix3 dwidenoise, on the full series before shell selection, and the denoised data given to all three toolkits. Denoising is thereby held constant and the estimator remains the only variable.

Giving each toolkit its own denoiser would not be a more faithful alternative. FSL provides none, so it would fit noisier data than the other two and any difference could no longer be attributed to the fit. Eddy-current correction was not tested; it corrects geometric distortion, whereas the finding under examination concerns eigenvalue estimation, which is driven by noise.

### The execution environment

Running three toolkits on byte-identical input is the practical obstacle to a comparison of this kind, and the reason it is rarely done. We removed it by building a single container in which all three are installed and driven from one interface, so the same data can be pushed through each without reinstallation, format conversion, or hand-managed paths. The container is what makes the comparison reproducible by others, and is released with the paper.

#### Container build

dMRI Rosetta Stone is implemented as a Streamlit web application (version ≥ 1.40; Streamlit Inc., 2019) and containerised with Docker (Merkel, 2014). Streamlit converts Python scripts into interactive browser-based interfaces without requiring HTML, CSS, or JavaScript, making the codebase accessible to any researcher with Python familiarity.

The Dockerfile uses a two-stage build. Stage 1 copies MRtrix3 3.0.4 binaries from the official mrtrix3/mrtrix3:latest image. Stage 2 starts from ubuntu:22.04, installs FSL 6.0.7 via the official fslinstaller.py script, copies MRtrix3 from Stage 1, and installs the Python stack — Streamlit, DIPY ≥ 1.7, nibabel, numpy, scipy, matplotlib, pandas, scikit-image — via pip. This multi-stage approach avoids dependency conflicts between FSL and MRtrix3 that arise from single-stage installation. The container serves the Streamlit interface at http://localhost:8501 via two commands:

docker build --platform linux/amd64 -t dmri-rosetta .
docker run --rm --platform linux/amd64 -p 8501:7860 dmri-rosetta

The first command builds the image (approximately 15 minutes, dominated by the FSL download); the second launches the container. No neuroimaging software is required on the host machine.

#### Pipeline coverage

The application covers seven pipeline stages, each as a dedicated page in the sidebar navigation (Table 1).

Brain Extraction. FSL bet, MRtrix3 dwi2mask, and DIPY median_otsu are demonstrated on the same b = 0 volume. Outputs are displayed as a red-tinted mask overlay on the b = 0 image (Figure 4).

Denoising (MP-PCA). Marchenko-Pastur PCA exploits diffusion signal redundancy to separate signal from thermal noise using random matrix theory (Veraart et al., 2016). MRtrix3 dwidenoise and DIPY mppca are demonstrated. FSL lacks a dedicated denoising tool; this asymmetry is made explicit with an explanation of when denoising is most beneficial.

Eddy Current and Motion Correction. FSL eddy_cpu (Andersson and Sotiropoulos, 2016), MRtrix3 dwifslpreproc (which wraps FSL eddy internally and prepends Gibbs-ringing removal via mrdegibbs), and DIPY motion_correction (rigid-body registration only, no eddy correction) are demonstrated. A demo subsample mode (5 b = 0 + 15 DWI volumes) reduces runtime from approximately 40 minutes to approximately 3 minutes for interactive use.

DTI Fitting. FSL dtifit, MRtrix3 dwi2tensor + tensor2metric, and DIPY TensorModel.fit() are applied to the same input data; FA, MD, AD, and RD maps are rendered for each toolkit (Figure 3).

Constrained Spherical Deconvolution. MRtrix3 dwi2fod (single-shell CSD, Tournier et al., 2007; multi-shell multi-tissue CSD, Jeurissen et al., 2014) and DIPY ConstrainedSphericalDeconvModel are demonstrated. FSL does not include CSD; its omission is discussed together with acquisition requirements for CSD (≥ 30 directions, high b-value).

Tractography. MRtrix3 tckgen with the iFOD2 probabilistic algorithm (Tournier et al., 2010) and streamline filtering via tcksift2 (Smith et al., 2015), FSL probtrackx2 (command display and conceptual overview), and DIPY LocalTracking (deterministic, DTI peaks) are demonstrated. Track density images (TDI) generated with tckmap provide voxelwise streamline density maps.

TBSS Voxelwise Group Analysis. The full TBSS pipeline (Smith et al., 2006) — tbss_1_preproc, tbss_2_reg, tbss_3_postreg, tbss_4_prestats, and randomise — is demonstrated. Because the platform operates on a single subject, a synthetic group is constructed by adding Gaussian noise realisations (σ = 0.03 FA units) to the real FA map. The absence of equivalent voxelwise pipelines in MRtrix3 and DIPY (fixel-based analysis in MRtrix3 is noted as the conceptual analogue) is made explicit.

#### Reference material

A dedicated Reference section provides: (1) a DTI metrics guide with the biological interpretation of FA, MD, AD, and RD and live map display; (2) a 16-term dMRI glossary; (3) a command cheat sheet covering all seven stages across all three toolkits; and (4) a decision framework for selecting the appropriate toolkit based on acquisition type, research question, and available compute resources.

#### Availability and requirements

| Project name | dMRI Rosetta Stone |
| --- | --- |
| Project home page | https://github.com/happybrotherhood/dmri-rosetta-stone |
| Archived version | v1.0.0, Zenodo, doi:10.5281/zenodo.22106455 (all versions: doi:10.5281/zenodo.22106454) |
| Operating system | Platform independent — Linux, macOS, and Windows, via Docker |
| Programming language | Python ≥ 3.9 |
| Other requirements | Docker ≥ 20.10. For local installation without Docker: FSL ≥ 6.0 and MRtrix3 ≥ 3.0.4 |
| Bundled toolkits | FSL 6.0.7, MRtrix3 3.0.4, DIPY ≥ 1.7 |
| Licence | MIT |
| Restrictions for non-academic use | None imposed by this project. Note, however, that the bundled FSL is distributed under the FSL Licence, which restricts commercial use; users intending commercial deployment must obtain a licence from Oxford University Innovation. MRtrix3 (MPL 2.0) and DIPY (BSD 3-clause) carry no such restriction. |

## Results

### The apparent difference between toolkits

Run as a user would normally run them — FSL dtifit --wls, MRtrix3 dwi2tensor at its default, DIPY fit_method="WLS" — the three toolkits do not agree equally well with one another (Table 2). MRtrix3 and DIPY were close on both datasets (FA r = 0.9990 and 0.9966), while FSL sat apart from both (FA r = 0.9604-0.9652 on Stanford, 0.8897-0.9157 on Sherbrooke). Read at face value this says FSL is the odd one out.

Bland-Altman analysis (Figure 3c) showed systematic offset rather than symmetric scatter in every FSL pairing. The offset did not behave consistently: FSL mean diffusivity was lower on both datasets, by 0.033 µm²/ms on Stanford and 0.127 on Sherbrooke, while the fractional anisotropy offset changed sign between them, FSL being lower on Stanford by 0.018 and higher on Sherbrooke by 0.014.

An offset that reverses sign between two ordinary acquisitions is difficult to explain as an implementation difference, and that inconsistency prompted the checks reported next.

### The toolkits were not running the same estimator

MRtrix3’s dwi2tensor does not perform plain weighted least squares by default. Its documentation states that it fits in two stages: weighted least squares using weights taken from the empirical signal intensities, followed by iteratively reweighted least squares in which the weights come from the signal predicted by the previous iteration, with two such iterations by default. FSL --wls and DIPY fit_method="WLS" perform a single weighted fit. The comparison above therefore contrasted two IWLS iterations against plain WLS, and attributed the result to the toolkit.

Repeating the comparison with MRtrix3 constrained to plain WLS (-iter 0) reverses the grouping entirely (Table 3). FSL and MRtrix3 then agree exactly: r = 1.0000 for both FA and MD on both datasets, with mean absolute error 0.0000 and a mean voxelwise difference of 1.1 × 10⁻⁴ FA units, which is numerical precision rather than agreement in the usual sense. MRtrix3 simultaneously moves away from DIPY, to r = 0.9602 on Stanford and 0.8895 on Sherbrooke — the same distance that had previously separated DIPY from FSL.

Two implementations of weighted least squares, written independently in different languages by different groups, return the same tensor. The apparent toolkit effect was an estimator effect.

### What separates DIPY is the source of the weights

DIPY remains apart from the other two even with MRtrix3 matched, and its documentation identifies why. Where MRtrix3’s first stage weights the fit by the measured signal, DIPY follows the two-pass scheme of Chung et al. (2006): an ordinary least-squares fit is performed first, and the signal it predicts supplies the weights for the weighted fit. FSL’s --wls agrees with MRtrix3 to numerical precision, so it uses measured-signal weights as well.

This divides the four arms into two families by weighting scheme rather than by software:

| Weights derived from | Arms |
| --- | --- |
| Measured signal | FSL dtifit --wls; MRtrix3 dwi2tensor -iter 0 |
| Predicted signal | DIPY fit_method="WLS"; MRtrix3 dwi2tensor default (iterated twice) |

The grouping predicts every correlation observed: within the measured-signal family r = 1.0000, within the predicted-signal family r = 0.9990, and between families r ≈ 0.96.

We tested this directly rather than resting on the correspondence. DIPY accepts user-supplied weights, so the same DIPY code was run twice on the same data, once with its default predicted-signal weights and once with measured-signal weights, and each compared against the FSL fit (Table 4). Supplying measured-signal weights reduced the FA disagreement with FSL from 0.0222 to 0.0003 mean absolute error on Stanford and from 0.0468 to 0.0035 on Sherbrooke; for MD the reduction was from 0.0320 to 0.0001 and from 0.1224 to 0.0022 µm²/ms. The bias fell to within 0.0023 FA units of zero in both datasets.

Changing one argument inside a single toolkit removes almost all of the difference previously attributed to the choice between toolkits.

### The same toolkit disagrees with itself more than with another

FSL dtifit performs ordinary least squares unless --wls is passed. The comparison so far used --wls, to match the estimator the other toolkits apply; the default is what most users run. Fitting the same data both ways places the size of the between-toolkit differences in context (Table 7).

Within FSL, switching between its default and --wls changed FA by a mean absolute 0.0176 on Stanford and 0.0505 on Sherbrooke. Between FSL --wls and MRtrix3 -iter 0 — different software, different language, different developers — the mean absolute difference was 0.0000 on both.

One flag inside one program therefore produced a larger difference than replacing the program entirely. The same holds against DIPY: FSL’s default sits 0.0191 and 0.0182 FA units from DIPY’s WLS on the two datasets, comparable to the between-scheme gap already reported, but reached without changing toolkit at all.

On the phantom FSL’s default recovers the truth about as well as the other arms, with FA bias +0.0028 against +0.0014 to +0.0022 for the weighted fits and a slightly higher RMSE of 0.0235 against 0.0209 to 0.0212 (Table 8). Ordinary least squares is a defensible estimator and is not being called incorrect here; the point is that it is a different one, selected by default, and that the choice is not recorded anywhere a reader of the methods section would see it.

### The effect is not confined to weighted least squares

Weighted least squares is one of several estimators the toolkits provide, and confining the comparison to it would leave open whether the pattern is an artefact of that choice. We therefore added DIPY’s non-linear fit, which works on the signal rather than its logarithm, and RESTORE, which downweights outliers, and compared every arm against the reference established above — FSL --wls, which is numerically identical to MRtrix3 -iter 0 (Table 7).

No estimator recovers the reference. On Stanford, mean absolute FA differences from it were 0.0219 (DIPY WLS), 0.0180 (NLLS), 0.0188 (RESTORE), 0.0244 (MRtrix3’s default IWLS) and 0.0173 (FSL’s default OLS); on Sherbrooke, 0.0465, 0.0376, 0.0392, 0.0451 and 0.0496. The spread across five distinct estimators, spanning three toolkits and both linear and non-linear fitting, is narrower than the gap between any of them and the measured-signal weighted fit.

This locates the effect more precisely than the initial contrast did. It is not that DIPY’s implementation is unusual; it is that one specific estimator — weighted least squares with weights taken from the measured signal — sits apart from every other option we tested, and two toolkits reach it only when explicitly asked.

### How large is the effect

Absolute differences in fractional anisotropy are hard to weigh without a scale. Expressed against the standard deviation of FA across white matter voxels in the same data, the differences reported above run from 0.12 to 0.25 standard deviations (Table 7).

The within-FSL comparison is the clearest case: changing one flag moved FA by 0.12 standard deviations on Stanford and 0.26 on Sherbrooke. A shift of a quarter of a standard deviation, produced by a command-line argument that is not recorded in any methods section we are aware of, is of the same order as effects that dMRI studies are designed to detect between groups.

### Accuracy depends on signal-to-noise ratio

That two estimators disagree does not by itself say either is inaccurate, and the question is settled only against a known truth. We fitted synthetic phantoms at three noise levels, since the theoretical objection to measured-signal weighting is that the weights are correlated with the noise they are meant to downweight, an effect that should grow as SNR falls (Table 7).

At SNR 30 the two families are close to the truth and to each other. Against a true FA of 0.7071 the measured-signal arms were biased by −0.0021 and the predicted-signal arms by +0.0014 to +0.0022, with root-mean-square errors of 0.021 in every arm. Read alone, this would suggest the choice does not matter.

It does at lower SNR. At SNR 20 the measured-signal bias grew to −0.0048 and at SNR 10 to −0.0177, against +0.0111 for DIPY’s two-pass scheme and +0.0182 for MRtrix3’s default. The mean diffusivity bias separated more sharply: −0.0408 µm²/ms for measured-signal weighting against +0.0172 for DIPY, a factor of 2.4. Root-mean-square error was higher for measured-signal weighting at every noise level tested.

The direction is consistent: measured-signal weighting underestimates both metrics and predicted-signal weighting overestimates them, with the former deviating further as noise increases. This reproduces, for the specific configurations these toolkits ship, the result Veraart et al. (2013) established in simulation.

It also explains a pattern in the real data. Differences between the families were roughly twice as large on Sherbrooke as on Stanford throughout, and Sherbrooke is the noisier acquisition — a single b = 0 volume against Stanford’s ten, and shells extending to b = 3500.

### Sensitivity to denoising

If the difference between weighting schemes is driven by noise, denoising should reduce it. MP-PCA denoising was applied once with MRtrix3 dwidenoise and the denoised series given to all three toolkits, leaving the estimator as the only variable (Table 8).

The differences shrink substantially but do not close. The FSL-to-MRtrix3-default MD offset fell from 0.033 to 0.019 µm²/ms on Stanford and from 0.127 to 0.049 on Sherbrooke, reductions of 44% and 61%. The FA offset fell correspondingly, and its sign still reversed between the two datasets, from −0.013 on Stanford to +0.003 on Sherbrooke. Agreement within the predicted-signal family was unaffected (MRtrix3 default against DIPY, r = 0.9990 before and after).

Roughly half of the between-family difference on real data is therefore attributable to noise, consistent with the phantom result that the two schemes differ little when noise is well behaved.

### Non-physical tensor fits

Unconstrained linear fitting can return negative eigenvalues, which produce fractional anisotropy above one and mean diffusivity below zero. Counts differ markedly between arms (Table 9). DIPY returned no such voxel in any run — two datasets, with and without denoising. FSL and MRtrix3 both returned them: on unprocessed data, 0.40% and 0.27% of white matter voxels on Stanford, and 3.97% and 2.03% on Sherbrooke.

Which of the two produces more is not stable. On unprocessed data FSL produced more than MRtrix3 in both datasets, but after denoising the order reversed on Stanford, MRtrix3 rising slightly to 0.29% while FSL fell to 0.23%. The robust statement is the one that held throughout: DIPY produced none, and the other two produced them at a rate that depends on both the acquisition and the preprocessing.

The rate tracks acquisition quality. Sherbrooke, with a single b = 0 volume against Stanford’s ten, produced roughly ten times as many failed fits in every arm that admits them.

These voxels barely move a robust statistic — excluding them changes mean absolute error by under 2% — but they dominate a correlation. Computed without a plausibility restriction, MRtrix3-DIPY MD agreement on Stanford appears to be r = 0.118 rather than 0.9965. Agreement statistics computed on raw tensor maps can therefore be wrong by an amount that changes the conclusion.

### Brain extraction

Brain masks produced by FSL bet, MRtrix3 dwi2mask and DIPY median_otsu, each run on the input its algorithm expects, overlapped at Dice coefficients of 0.9009 to 0.9277 on Stanford and 0.9058 to 0.9668 on Sherbrooke (Table 2), with no pair below 0.90.

Pairwise Dice conceals how differently the algorithms behave. On Stanford the mask volumes were 203,984 voxels (FSL), 187,948 (DIPY) and 167,950 (MRtrix3), a spread of 21% between largest and smallest despite every pairwise coefficient exceeding 0.90. The ordering is not stable either: on Sherbrooke the volumes were 199,249 (MRtrix3), 188,717 (FSL) and 186,450 (DIPY), a spread of 7%, with MRtrix3 moving from the most conservative mask on one dataset to the most inclusive on the other.

Disagreements concentrate at the cortical boundary and, distinctively, at the lateral ventricles (Figure 4), which DIPY median_otsu excludes and the other two retain. A mask that includes ventricles admits high-diffusivity, near-isotropic voxels into any subsequent group statistics.

## Discussion

### What the toolkit label does and does not tell you

The result that organises the others is that FSL and MRtrix3, given the same estimator, return the same tensor to numerical precision. Two independent implementations — different languages, different groups, different decades — agree at r = 1.0000 with a mean voxelwise difference of about 10⁻⁴ FA units. Implementation quality is not the variable.

What varies is which estimator each toolkit applies when the user does not specify one. MRtrix3 ships two iterations of reweighted least squares; FSL performs ordinary least squares unless --wls is passed; DIPY performs a single weighted fit but derives the weights from a preliminary OLS prediction rather than from the measured signal. None of this appears in the command a user types. dwi2tensor dwi.mif tensor.mif and dtifit -k data -o out ... look like the same operation described in two dialects, which is how the field generally treats them, and they are not.

The comparison within FSL makes the size of this plain. Changing one flag in one program moved FA by 0.0176 and 0.0505 mean absolute units on the two datasets, while replacing that program with MRtrix3 at a matched estimator moved it by 0.0000. Whatever “software-related variability” denotes, it is not a property of the software.

Expressed on a scale that can be weighed, those shifts are 0.12 and 0.26 standard deviations of white matter FA. That is not a rounding difference. It is the same order as the group effects dMRI studies are powered to detect, produced by an argument that appears in no methods section we have read. Widening the comparison to five estimators across three toolkits, linear and non-linear, does not change the picture: every one of them sits apart from the measured-signal weighted fit, and the spread among them is smaller than their common distance from it.

This reframes what “software-related variability” means in dMRI. Reported differences between toolkits are real, and our first analysis reproduced them, but they are not differences between codebases. They are differences between statistical estimators that happen to be bundled as defaults. The distinction matters because the two have different remedies: a codebase difference would require the developers to act, whereas an estimator difference can be removed by the analyst, in our case with a single command-line argument.

### The literature has an answer; the toolkits have not adopted it

Veraart et al. (2013) reported that weighting by the measured signal degrades accuracy and that multi-step weighting performs better. Our phantoms reproduce this for the configurations these toolkits ship, and locate where it matters: at SNR 30 the choice is close to irrelevant, while at SNR 10 the measured-signal scheme carries roughly twice the mean diffusivity bias and a higher root-mean-square error.

A decade later the toolkits have not converged. MRtrix3 applies multi-step reweighting by default, and DIPY the two-pass scheme of Chung et al. (2006), so both follow the recommendation without the user needing to know it exists. FSL performs ordinary least squares by default, and its --wls option implements the single-pass measured-signal weighting that Veraart et al. specifically caution against. A user who reads that --wls enables weighted least squares, and reasonably concludes this is the more principled option, gets the scheme the literature advises against.

We are not proposing a new estimator or contradicting the existing recommendation. The contribution is to show that the recommendation has not reached the defaults, that the resulting differences are of a size that matters at realistic SNR, and that a user cannot see which scheme they are applying from the command they run.

### The reporting problem

Methods sections routinely record the toolkit and version. On the evidence here that is the wrong level of description. Two studies both reporting “MRtrix3 3.0.4” may have used different estimators if one passed -iter 0; two studies reporting different toolkits may have used the same one. The toolkit name is neither necessary nor sufficient to reconstruct what was computed.

What would suffice is short: the estimator, the weighting scheme, and any non-default fitting arguments. We would encourage journals and reporting checklists to ask for it, and we note that our own first analysis — which compared IWLS against WLS and concluded that one toolkit was aberrant — is an example of what happens when this information is absent.

### Non-physical fits as a quality indicator

DIPY returned no physically impossible tensor in any configuration tested, while FSL and MRtrix3 returned them at rates between 0.14% and 3.97% depending on acquisition and preprocessing. We report this without proposing a mechanism, and note that which of the two produces more is not stable: denoising reversed their order on one dataset.

The practical value lies less in ranking toolkits than in the diagnostic. The count of inadmissible voxels responds to acquisition quality — Sherbrooke’s single b = 0 volume produced roughly ten times the rate of Stanford’s ten — and is cheap to compute. It is also a trap: left in place these voxels reduced an apparent MD correlation from 0.997 to 0.118 while moving the mean absolute error by under 2%. Agreement statistics on tensor-derived maps should be computed only over voxels whose values are physically admissible, and the excluded count reported.

### Limitations

The real-data component rests on one subject from each of two datasets. This is sufficient for the central claim, which concerns whether two implementations of one estimator agree — a question that does not require a sample, since the answer was exact agreement. It is not sufficient to characterise how the between-scheme difference varies with acquisition, and we do not attempt to.

The comparison covers the tensor fit under default and matched settings, on unprocessed and denoised data. Eddy-current and motion correction were not applied, so the numbers describe two operating points rather than a complete pipeline. Only FA and MD were examined; the analysis does not extend to models with more parameters, where the scope for divergence is larger.

The phantom is a deliberately simple test: Gaussian noise, no artefacts, no partial volume, no motion. It establishes that neither weighting scheme is systematically inaccurate under favourable conditions. It cannot establish which behaves better under unfavourable ones, which is the question the real-data gap raises and which we leave open.

### Extensions

Adding datasets is cheap by design — one fetcher entry and two commands — and multi-subject collections would allow the between-scheme difference to be characterised against acquisition parameters rather than merely observed at two points. The same controlled arrangement applies to the preprocessing steps excluded here, taken one at a time, and to models beyond the tensor, where implementations diverge more. Each toolkit also exposes fitting options left at their defaults here; whether the between-scheme gap persists across them is directly testable.

## Conclusion

Given identical input, FSL and MRtrix3 constrained to the same estimator return the same diffusion tensor to numerical precision. The differences reported between diffusion MRI toolkits are not differences between implementations. They are differences between statistical estimators, distributed as defaults that the user’s command does not reveal: MRtrix3 iterates its reweighting twice, DIPY derives its weights from a predicted rather than a measured signal, and FSL performs ordinary least squares unless asked otherwise.

Which estimator to prefer is not an open question. Veraart et al. (2013) showed that measured-signal weighting degrades accuracy and that multi-step weighting performs better, and our phantoms reproduce this for the shipped configurations: negligible at SNR 30, but at SNR 10 the measured-signal scheme carries roughly twice the mean diffusivity bias. A decade on, MRtrix3 and DIPY follow the recommendation by default while FSL’s weighted option does not, and nothing in the interface tells the user which they have.

The practical consequence is a reporting one. Recording the toolkit and version does not describe what was computed; recording the estimator and its settings does, and costs one line. Our own first pass through this comparison concluded that one toolkit was aberrant, on evidence that turned out to reflect an unmatched default — which is the clearest argument we can offer for making that line standard.

## CRediT authorship contribution statement

Busra Mutlu: Conceptualization, Methodology, Software, Formal analysis, Investigation, Data curation, Visualization, Writing - original draft, Writing - review and editing.

## Declaration of competing interest

The author declares that she has no known competing financial interests or personal relationships that could have appeared to influence the work reported in this paper.

## Funding

[TO COMPLETE: name the funder and grant number, or state that this research received no specific grant from funding agencies in the public, commercial, or not-for-profit sectors.]

## Declaration of generative AI and AI-assisted technologies in the writing process

[TO COMPLETE OR DELETE — see the note accompanying this manuscript. If generative AI was used to assist with drafting or editing, Elsevier requires a statement here naming the tool and describing its use, and confirming that the author reviewed and edited the output and takes full responsibility for the content of the publication.]

## Data availability

The complete source code, Dockerfile, and documentation are openly available at https://github.com/happybrotherhood/dmri-rosetta-stone under an MIT licence. The exact version reported here is archived on Zenodo as v1.0.0, doi:10.5281/zenodo.22106455; the concept DOI doi:10.5281/zenodo.22106454 always resolves to the most recent version. Both datasets analysed — Stanford HARDI and Sherbrooke 3-shell — are distributed openly by the DIPY project at https://dipy.org and are retrieved automatically by scripts/fetch_sample_data.py. No registration or credentials are required to reproduce any result presented here, and the four commands that regenerate every reported value are given at the end of the Materials and Methods.

## Acknowledgements

[TO COMPLETE OR DELETE: supervisors, colleagues, computing resources.]

## References

Avesani P, McPherson B, Hayashi S, Caiafa CF, Henschel R, Garyfallidis E, et al. (2019) The open diffusion data derivatives, brain data upcycling via integrated collection and reuse. Scientific Data 6, 69. doi: 10.1038/s41597-019-0073-y

Andersson JLR and Sotiropoulos SN (2016) An integrated approach to correction for off-resonance effects and subject movement in diffusion MR imaging. NeuroImage 125, 1063–1078. doi: 10.1016/j.neuroimage.2015.10.019

Basser PJ, Mattiello J, and LeBihan D (1994) MR diffusion tensor spectroscopy and imaging. Biophysical Journal 66, 259–267. doi: 10.1016/S0006-3495(94)80775-1

Beaulieu C (2002) The basis of anisotropic water diffusion in the nervous system – a technical review. NMR in Biomedicine 15, 435–455. doi: 10.1002/nbm.782

Bhagwat N, Barry A, Dickie EW, Brown ST, Devenyi GA, Hatano K, et al. (2021) Understanding the impact of preprocessing pipelines on neuroimaging cortical surface analyses. GigaScience 10, giaa155. doi: 10.1093/gigascience/giaa155

Catani M and Thiebaut de Schotten M (2008) A diffusion tensor imaging tractography atlas for virtual in vivo dissections. Cortex 44, 1105–1132. doi: 10.1016/j.cortex.2008.05.004

Chung S, Lu Y, and Henry RG (2006) Comparison of bootstrap approaches for estimation of uncertainties of DTI parameters. NeuroImage 33, 531–541. doi: 10.1016/j.neuroimage.2006.07.001

Garyfallidis E, Brett M, Amirbekian B, Rokem A, Van Der Walt S, Descoteaux M, et al. (2014) Dipy, a library for the analysis of diffusion MRI data. Frontiers in Neuroinformatics 8, 8. doi: 10.3389/fninf.2014.00008

Glasser MF, Sotiropoulos SN, Wilson JA, Coalson TS, Fischl B, Andersson JL, et al. (2013) The minimal preprocessing pipelines for the Human Connectome Project. NeuroImage 80, 105–124. doi: 10.1016/j.neuroimage.2013.04.127

Jahn A (2020) Andy’s Brain Book: An Introduction to Neuroimaging Analysis. Available at: https://andysbrainbook.readthedocs.io (Accessed June 2026).

Jenkinson M, Beckmann CF, Behrens TEJ, Woolrich MW, and Smith SM (2012) FSL. NeuroImage 62, 782–790. doi: 10.1016/j.neuroimage.2011.09.015

Jeurissen B, Tournier JD, Dhollander T, Connelly A, and Sijbers J (2014) Multi-tissue constrained spherical deconvolution for improved analysis of multi-shell diffusion MRI data. NeuroImage 103, 411–426. doi: 10.1016/j.neuroimage.2014.07.061

Jones DK (ed.) (2010) Diffusion MRI: Theory, Methods, and Applications. Oxford: Oxford University Press.

Koay CG, Chang LC, Carew JD, Pierpaoli C, and Basser PJ (2006) A unifying theoretical and algorithmic framework for least squares methods of estimation in diffusion tensor imaging. Journal of Magnetic Resonance 182, 115–125. doi: 10.1016/j.jmr.2006.06.020

Merkel D (2014) Docker: Lightweight Linux containers for consistent development and deployment. Linux Journal 2014, 2.

Richie-Halford A, Cieslak M, Ai L, Caffarra S, Covitz S, Franco AR, et al. (2022) An analysis-ready and quality controlled resource for pediatric brain white-matter research. Scientific Data 9, 616. doi: 10.1038/s41597-022-01695-7

Rokem A, Yeatman JD, Pestilli F, Kay KN, Mezer A, van der Walt S, et al. (2015) Evaluating the accuracy of diffusion MRI models in white matter. PLOS ONE 10, e0123272. doi: 10.1371/journal.pone.0123272

Smith RE, Tournier JD, Calamante F, and Connelly A (2015) SIFT2: Enabling dense quantitative assessment of brain white matter connectivity using streamlines tractography. NeuroImage 119, 338–351. doi: 10.1016/j.neuroimage.2015.06.092

Smith SM, Jenkinson M, Woolrich MW, Beckmann CF, Behrens TEJ, Johansen-Berg H, et al. (2004) Advances in functional and structural MR image analysis and implementation as FSL. NeuroImage 23 (Suppl 1), S208–S219. doi: 10.1016/j.neuroimage.2004.07.051

Smith SM, Jenkinson M, Johansen-Berg H, Rueckert D, Nichols TE, Mackay CE, et al. (2006) Tract-based spatial statistics: Voxelwise analysis of multi-subject diffusion data. NeuroImage 31, 1487–1505. doi: 10.1016/j.neuroimage.2006.02.024

Streamlit Inc (2019) Streamlit: The fastest way to build and share data apps [Computer software]. Available at: https://streamlit.io

Tournier JD, Calamante F, and Connelly A (2007) Robust determination of the fibre orientation distribution in diffusion MRI: Non-negativity constrained super-resolved spherical deconvolution. NeuroImage 35, 1459–1472. doi: 10.1016/j.neuroimage.2007.02.016

Tournier JD, Calamante F, and Connelly A (2010) Improved probabilistic streamlines tractography by 2nd order integration over fibre orientation distributions. Proceedings of the International Society for Magnetic Resonance in Medicine 18, 1670.

Tournier JD, Smith RE, Raffelt D, Tabbara R, Dhollander T, Pietsch M, et al. (2019) MRtrix3: A fast, flexible and open software framework for medical image processing and visualisation. NeuroImage 202, 116137. doi: 10.1016/j.neuroimage.2019.116137

Van Essen DC, Smith SM, Barch DM, Behrens TEJ, Yacoub E, Ugurbil K, et al. (2013) The WU-Minn Human Connectome Project: An overview. NeuroImage 80, 62–79. doi: 10.1016/j.neuroimage.2013.05.041

Veraart J, Sijbers J, Sunaert S, Leemans A, and Jeurissen B (2013) Weighted linear least squares estimation of diffusion MRI parameters: strengths, limitations, and pitfalls. NeuroImage 81, 335–346. doi: 10.1016/j.neuroimage.2013.05.028

Veraart J, Novikov DS, Christiaens D, Ades-aron B, Sijbers J, and Fieremans E (2016) Denoising of diffusion MRI using random matrix theory. NeuroImage 142, 394–406. doi: 10.1016/j.neuroimage.2016.08.016

## Tables

Table 1. Tensor-fitting configurations compared. The first three rows are each toolkit as a user would normally invoke it; the fourth places MRtrix3 on the same estimator as the other two. Weighting scheme as documented by each project.

| Arm | Command | Estimator | Weights derived from | Is it the default? |
| --- | --- | --- | --- | --- |
| FSL, default | dtifit | OLS | unweighted | yes |
| FSL, weighted | dtifit --wls | WLS | measured signal | no |
| MRtrix3, default | dwi2tensor | WLS + 2 IWLS | predicted signal | yes |
| MRtrix3, matched | dwi2tensor -iter 0 | WLS | measured signal | no |
| DIPY | TensorModel(fit_method="WLS") | WLS | predicted signal, from an initial OLS fit | yes |

Table 2. Toolkits at their default settings. Agreement over white matter voxels (FA > 0.2 in all three), with every input held identical. Pearson r, mean absolute error (MAE) and Bland-Altman mean bias. MD in µm²/ms. Statistics use only physically admissible voxels (FA in [0, 1]; 0 < MD ≤ 3.0 × 10⁻³ mm²/s). Mask DSC compares each pair of brain extractions, each tool run on the input its algorithm expects.

Stanford HARDI (n = 65,002 white matter voxels)

| Comparison | Mask DSC | FA r | FA MAE | FA bias | MD r | MD MAE | MD bias |
| --- | --- | --- | --- | --- | --- | --- | --- |
| FSL vs. MRtrix3 | 0.9024 | 0.9652 | 0.0247 | -0.0181 | 0.9402 | 0.0333 | -0.0330 |
| FSL vs. DIPY | 0.9277 | 0.9604 | 0.0228 | -0.0158 | 0.9368 | 0.0325 | -0.0322 |
| MRtrix3 vs. DIPY | 0.9009 | 0.9990 | 0.0029 | +0.0027 | 0.9965 | 0.0012 | +0.0009 |

Sherbrooke 3-shell (n = 111,032 white matter voxels)

| Comparison | Mask DSC | FA r | FA MAE | FA bias | MD r | MD MAE | MD bias |
| --- | --- | --- | --- | --- | --- | --- | --- |
| FSL vs. MRtrix3 | 0.9058 | 0.9157 | 0.0460 | +0.0139 | 0.9020 | 0.1271 | -0.1271 |
| FSL vs. DIPY | 0.9668 | 0.8897 | 0.0504 | +0.0148 | 0.9012 | 0.1272 | -0.1272 |
| MRtrix3 vs. DIPY | 0.9163 | 0.9966 | 0.0059 | +0.0053 | 0.9991 | 0.0029 | -0.0009 |

Table 3. The same comparison with the estimator matched. MRtrix3 constrained to plain weighted least squares (dwi2tensor -iter 0), alongside FSL dtifit --wls and DIPY fit_method="WLS". Compare with Table 2, where MRtrix3 ran its default of two reweighting iterations.

Stanford HARDI

| Comparison | FA r | FA MAE | FA bias | MD r | MD MAE | MD bias |
| --- | --- | --- | --- | --- | --- | --- |
| FSL vs. MRtrix3 | 1.0000 | 0.0000 | +0.0000 | 1.0000 | 0.0000 | +0.0000 |
| FSL vs. DIPY | 0.9601 | 0.0228 | -0.0158 | 0.9367 | 0.0325 | -0.0322 |
| MRtrix3 vs. DIPY | 0.9602 | 0.0228 | -0.0158 | 0.9366 | 0.0325 | -0.0322 |

Sherbrooke 3-shell

| Comparison | FA r | FA MAE | FA bias | MD r | MD MAE | MD bias |
| --- | --- | --- | --- | --- | --- | --- |
| FSL vs. MRtrix3 | 1.0000 | 0.0000 | +0.0000 | 1.0000 | 0.0000 | +0.0000 |
| FSL vs. DIPY | 0.8895 | 0.0505 | +0.0149 | 0.9013 | 0.1274 | -0.1274 |
| MRtrix3 vs. DIPY | 0.8895 | 0.0505 | +0.0149 | 0.9012 | 0.1274 | -0.1274 |

Table 4. Manipulating the weighting scheme inside one toolkit. The same DIPY code run twice on the same data, once with its default weights (squared signal predicted by an initial OLS fit) and once with measured-signal weights, each compared against the FSL dtifit --wls fit. MD in µm²/ms.

| Dataset | DIPY weights | FA r | FA MAE | FA bias | MD r | MD MAE | MD bias |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stanford | predicted (default) | 0.9662 | 0.0222 | +0.0152 | 0.9579 | 0.0320 | +0.0317 |
|  | measured (as FSL) | 0.9989 | 0.0003 | -0.0001 | 0.9982 | 0.0001 | +0.0001 |
| Sherbrooke | predicted (default) | 0.9173 | 0.0468 | -0.0201 | 0.9136 | 0.1224 | +0.1224 |
|  | measured (as FSL) | 0.9949 | 0.0035 | -0.0023 | 0.9962 | 0.0022 | +0.0022 |

Table 5. The same toolkit against itself. FSL dtifit at its default (ordinary least squares) compared with FSL --wls, and with the matched-estimator fits from the other toolkits. White matter voxels admissible in every arm.

| Dataset | Comparison | FA r | FA MAE | FA bias |
| --- | --- | --- | --- | --- |
| Stanford | FSL --wls vs. FSL default (OLS) | 0.9649 | 0.0176 | +0.0016 |
|  | FSL --wls vs. MRtrix3 -iter 0 | 1.0000 | 0.0000 | −0.0000 |
|  | FSL default (OLS) vs. DIPY WLS | 0.9921 | 0.0191 | −0.0171 |
| Sherbrooke | FSL --wls vs. FSL default (OLS) | 0.9069 | 0.0505 | +0.0166 |
|  | FSL --wls vs. MRtrix3 -iter 0 | 1.0000 | 0.0000 | +0.0000 |
|  | FSL default (OLS) vs. DIPY WLS | 0.9884 | 0.0182 | +0.0013 |

Table 6. Every estimator against the measured-signal weighted fit. The reference is FSL dtifit --wls, which is numerically identical to MRtrix3 dwi2tensor -iter 0. MAE/SD expresses the mean absolute difference as a fraction of the standard deviation of FA across white matter voxels in the same data, so that the size of each difference can be weighed. Arms marked (default) are what the toolkit does without options.

| Dataset | Arm | FA r | FA MAE | FA bias | MAE / SD |
| --- | --- | --- | --- | --- | --- |
| Stanford | FSL default, OLS (default) | 0.9676 | 0.0173 | +0.0015 | 0.12 |
|  | MRtrix3, IWLS ×2 (default) | 0.9683 | 0.0244 | -0.0184 | 0.17 |
|  | DIPY WLS (default) | 0.9686 | 0.0219 | -0.0156 | 0.15 |
|  | DIPY NLLS | 0.9781 | 0.0180 | -0.0125 | 0.12 |
|  | DIPY RESTORE | 0.9730 | 0.0188 | -0.0128 | 0.13 |
| Sherbrooke | FSL default, OLS (default) | 0.9126 | 0.0496 | +0.0171 | 0.25 |
|  | MRtrix3, IWLS ×2 (default) | 0.9213 | 0.0451 | +0.0142 | 0.23 |
|  | DIPY WLS (default) | 0.9176 | 0.0465 | +0.0186 | 0.24 |
|  | DIPY NLLS | 0.9386 | 0.0376 | +0.0178 | 0.19 |
|  | DIPY RESTORE | 0.9344 | 0.0392 | +0.0165 | 0.20 |

Table 7. Accuracy against a known ground truth, as a function of noise. Synthetic phantom, single-fibre region, true FA = 0.7071 and true MD = 0.7000 µm²/ms. Arms are grouped by the source of their weights. The theoretical objection to measured-signal weighting is that the weights are correlated with the noise, so the difference between families is expected to grow as SNR falls (Veraart et al., 2013).

| SNR | Arm | Weights | FA bias | FA RMSE | MD bias | MD RMSE |
| --- | --- | --- | --- | --- | --- | --- |
| 30 | FSL –wls | measured | -0.0021 | 0.0212 | -0.0050 | 0.0180 |
|  | MRtrix3 -iter 0 | measured | -0.0021 | 0.0212 | -0.0050 | 0.0180 |
|  | DIPY WLS | predicted | +0.0014 | 0.0209 | +0.0017 | 0.0174 |
|  | MRtrix3 default (IWLS) | predicted | +0.0022 | 0.0209 | +0.0020 | 0.0175 |
| 20 | FSL –wls | measured | -0.0048 | 0.0322 | -0.0110 | 0.0281 |
|  | MRtrix3 -iter 0 | measured | -0.0048 | 0.0322 | -0.0110 | 0.0281 |
|  | DIPY WLS | predicted | +0.0030 | 0.0312 | +0.0039 | 0.0265 |
|  | MRtrix3 default (IWLS) | predicted | +0.0048 | 0.0314 | +0.0047 | 0.0267 |
| 10 | FSL –wls | measured | -0.0177 | 0.0682 | -0.0408 | 0.0653 |
|  | MRtrix3 -iter 0 | measured | -0.0177 | 0.0682 | -0.0408 | 0.0653 |
|  | DIPY WLS | predicted | +0.0111 | 0.0623 | +0.0172 | 0.0575 |
|  | MRtrix3 default (IWLS) | predicted | +0.0182 | 0.0637 | +0.0215 | 0.0598 |

Table 8. Sensitivity to denoising. The same comparison after MP-PCA denoising applied once and given to all three toolkits. Toolkits at their default settings, as in Table 2. MD in µm²/ms.

| Dataset | Comparison | FA bias raw | FA bias denoised | MD bias raw | MD bias denoised |
| --- | --- | --- | --- | --- | --- |
| Stanford | FSL vs. MRtrix3 | -0.0181 | -0.0134 | -0.0330 | -0.0186 |
|  | FSL vs. DIPY | -0.0158 | -0.0110 | -0.0322 | -0.0178 |
|  | MRtrix3 vs. DIPY | +0.0027 | +0.0026 | +0.0009 | +0.0009 |
| Sherbrooke | FSL vs. MRtrix3 | +0.0139 | +0.0034 | -0.1271 | -0.0494 |
|  | FSL vs. DIPY | +0.0148 | +0.0048 | -0.1272 | -0.0502 |
|  | MRtrix3 vs. DIPY | +0.0053 | +0.0031 | -0.0009 | -0.0011 |

Table 9. Non-physical tensor fits within the white matter mask. Voxels violating the definition of each metric, with percentage of the mask in parentheses, before and after denoising.

| Dataset | Metric | Preprocessing | FSL | MRtrix3 | DIPY |
| --- | --- | --- | --- | --- | --- |
| Stanford | FA > 1 | none | 261 (0.40%) | 178 (0.27%) | 0 (0.00%) |
| Stanford | MD ≤ 0 | none | 261 (0.40%) | 91 (0.14%) | 0 (0.00%) |
| Stanford | FA > 1 | MP-PCA | 154 (0.23%) | 193 (0.29%) | 0 (0.00%) |
| Stanford | MD ≤ 0 | MP-PCA | 105 (0.16%) | 101 (0.15%) | 0 (0.00%) |
| Sherbrooke | FA > 1 | none | 4,404 (3.97%) | 2,252 (2.03%) | 0 (0.00%) |
| Sherbrooke | MD ≤ 0 | none | 3,016 (2.72%) | 603 (0.54%) | 0 (0.00%) |
| Sherbrooke | FA > 1 | MP-PCA | 2,675 (2.56%) | 2,092 (2.00%) | 0 (0.00%) |
| Sherbrooke | MD ≤ 0 | MP-PCA | 1,330 (1.27%) | 641 (0.61%) | 0 (0.00%) |

## Figure Captions

Figure 1. Design of the comparison. Both datasets are open and need no credentials. Left: brain extraction is compared with each tool run on the input its algorithm expects, so the Dice coefficients reflect the difference a user would meet in practice. Right: tensor fitting is compared with every arm given identical input — one brain mask, one volume subset of b = 0 plus a single non-zero shell, no preprocessing — and the estimator varied deliberately instead. Arms are labelled by the source of their weights: FSL --wls and MRtrix3 -iter 0 weight by the measured signal, while DIPY and MRtrix3’s default weight by a predicted signal. Statistics are then restricted to physically admissible voxels, without which a few hundred failed fits dominate the correlation, and every arm is checked against a phantom with known eigenvalues. Toolkit colours are consistent across Figures 1, 3 and 4. Generated by scripts/make_fig1_design.py.

Figure 2. The environment at the tensor-fitting stage, on the Stanford HARDI dataset. The sidebar reports the resolved dataset — detected shells and volume count — and lists the pipeline pages; a collapsible panel reports which toolkit binaries are present. The main panel presents one tab per toolkit, and the selected tab shows the exact command, the outputs it will produce, a Run button that submits it via Python subprocess and streams output live, and the resulting maps. Commands are shown in full and with relative paths, so any of them can be copied and run outside the environment unchanged.

Figure 3. Agreement between toolkits at their default settings, on the Stanford HARDI dataset; the Sherbrooke equivalent is Supplementary Figure S1. (a) FA maps from FSL dtifit, MRtrix3 tensor2metric and DIPY TensorModel, applied to identical input with one shared brain mask. (b) Voxelwise FA scatter for each pair over white matter voxels (FA > 0.2 in all three), with Pearson r, mean absolute error and identity line. (c) Bland–Altman plots showing voxelwise FA differences against their mean; horizontal lines mark the mean difference and ±1.96 SD. The MRtrix3–DIPY pair clusters tightly about the identity line while both FSL pairings show systematic offset — a pattern that Table 3 shows to follow the estimator rather than the toolkit. Voxels with physically inadmissible values are excluded.

Figure 4. Brain extraction on the Stanford HARDI dataset; the Sherbrooke equivalent is Supplementary Figure S2. The b = 0 mean image in greyscale with the brain mask overlaid in red, from FSL bet (left, 203,984 voxels), MRtrix3 dwi2mask (centre, 167,950) and DIPY median_otsu (right, 187,948). Dice coefficients are in Table 2. Two differences are visible: the extent of cortical boundary coverage, and the treatment of the lateral ventricles, which median_otsu excludes and the other two retain.

Supplementary Figure S1. As Figure 3, for the Sherbrooke 3-shell dataset (b = 0 + b = 1000 s/mm² subset).

Supplementary Figure S2. As Figure 4, for the Sherbrooke 3-shell dataset.